## Slide 1
INTERNATIONAL DEPLOYMENT OF ACCELEROMETERS
Peter Davis, IGPP
Quality Control of IRIS/IDA data using MUSTANG
Managing Data from Seismic Networks Workshop
Hanoi, Vietnam
September 2015

## Slide 3: Telemetry links:
LAN
leased lines / VSATs
local ISP

## Slide 4
 Tools to monitor flow
Robust system to telemeter data

## Slide 5
Timing
Waveform quality
Accuracy of instrument responses (metadata)
At GSN stations, QC exploits presence of more than one broadband sensor
Past experience guides what problems to look for
Quality control focuses on:

## Slide 6
DCC staff checks for proper clock operation and for mass position offsets
sample rate correction (clocks on old systems are permitted to drift freely)
Verify Timing

## Slide 7
Timing check:
(We make use of earthquake locations announced by USGS NEIC.)
Quake 1
Quake 2
Quake 3

## Slide 8
compare primary and secondary sensors (We exploit data redundancy.)
KS-54000 – OK
STS2 – failed

## Slide 9: Seismometer masses saturated

## Slide 10
problems unique to tri-axial sensors
STS2
STS2
STS2
KS-54000
KS-54000
KS-54000

## Slide 11
loss of vacuum (We examine long data segments.)

## Slide 12: Excessive Spiking:
Clipping

## Slide 13: Take advantage of IRIS resources:
Web services: http://service.iris.edu/
MUSTANG: http://service.iris.edu/mustang/

## Slide 14: II.ARU:
Equipment was relocated to new vault by EME staff
Was the new vault quieter?

## Slide 15: MUSTANG example:
http://service.iris.edu/mustang/noise-pdf/1/query?net=II&sta=ARU&loc=00&cha=BHN&quality=M&starttime=2010-01-01&endtime=2010-01-31&format=plot&plot.title=II.ARU.BHN&plot.interpolation=bicubic
Before relocation
After relocation
http://service.iris.edu/mustang/noise-pdf/docs/1/builder/

## Slide 16
Checking performance of strong motion instruments
Replacing legacy code with platform-independent routines
Using synthetic seismogram resources
Using Web Services for QC:

## Slide 17
Mixture of sensor models (FBA-23, Episensor) and DASs (Reftek, SAN, Q330) in network; signal detection threshold varies widely
Most signals below detection threshold for these sensors – hard to evaluate if unit is operating correctly
Strong motion sensor issues:

## Slide 18
A well performing Episensor :
Signal too small for FBA23
Blue=broadband sensor Red=Episensor/FBA
Used web services product IrisFetch.m to (1) find earthquake, (2) retrieve data into Matlab©, (3) deconvolve signal, then (4) plot data.

## Slide 19
A plot produced by a very old computer program…
Replacing legacy code:

## Slide 20
P-wave arrivals for both the primary (blue) and secondary (red) sensors at II.PFO. Plot replaces legacy code using web services under Matlab©.
Used web services product IrisFetch.m to (1) find earthquake, (2) retrieve data into Matlab©, (3) deconvolve signal, then (4) plot data.

## Slide 21: and for waveform quality:
Used web services product IrisFetch.m to (1) find earthquake, (2) retrieve data into Matlab©, (3) deconvolve signal, then (4) plot data.

## Slide 22
Match of Specfem3D synthetics (red) to observed data (black) for event at II.PFO. (Web services)
(Synthetics available courtesy J. Tromp, Princeton University.)
Using synthetic seismograms

## Slide 23: Use of MUSTANG in analysis of WRAB
http://service.iris.edu/mustang/measurements/docs/1/builder/
“Transfer function metric”

## Slide 24
  Computed at IRIS DMC
  Reviewed at IDA DCC
  Other examples include:
    percent_availability
    Mean
    Median
    RMS
    SNR
    up_down_times
Example of MUSTANG metric

## Slide 25
IDA leverages academic research results to improve quality control:
Lamont Waveform Quality Center’s measurements are reviewed regularly for evidence of problems with sensor response information.

## Slide 26
Checking that the metadata are correct…

## Slide 27
Traditional calibration procedures verify the DAS sensitivity and measure shape of system response.
Metadata Accuracy

## Slide 28
DC Calibration Test

## Slide 29: Random binary calibration tests are performed annually to determine the shape of the sensor’s response.
Calibration

## Slide 30
Convolve model of sensor response with input
Compare observed output with model
Iterate until fit acceptable

## Slide 31
Individual elements of recording systems are well calibrated, but overall system response may be in error.
Independent recording system used to check the GSN system in place with no need to perturb in situ instruments.
Seismometer calibration:

## Slide 32: Calibrate with independent, portable sensor:
Nanometrics compact Trillium 120 and Taurus digitizer

## Slide 33
APS used to measure orientation of Trillium
Vault at II- ESK (Eskdalemuir, Scoltland)

## Slide 34
Accurate determination of relative response of the common instruments
Accurate determination of relative orientation
2+ hours of microseismic background noise is sufficient for accurate reading

## Slide 35
Reference sensor
Reference sensor
GSN sensor
GSN sensor
Sensor Calibration and Orientation

## Slide 36
Analysis of microseisms yield relative amplitude accurate to < 1%, orientation < 1 deg

## Slide 37
Shaketable at UCSD

## Slide 38
Effectiveness of the calibration program can be assessed using long period modes excited by large earthquakes
First used by Davis et al. (2005)

## Slide 39
Sumatra-Andaman earthquake excited many low frequency normal modes
Figure from Park et al. (2004), courtesy G. Roult

## Slide 40
Distributions identify problematic responses…

## Slide 41
Broad adoption of FDSN web services standards leads to many useful quality assessment tools
  Checking performance of strong motion sensors
  Replacing legacy code with platform-independent routines
  Using synthetic seismogram resources
IRIS MUSTANG program promoting improved quality control of GSN and other data sets
Software Developments:

## Slide 42
Pier at the Scripps Institution of Oceanography, UCSD